## LHO End Y
Laser SN08 | Incoming laser power from T1300100
Date | 2013-06-04 | Aug. 3, 2018 | Mar. 19, 2019 | Jan, 23, 2024 | May , 21,2025 | Feb 3rd 2026
Laser Shutter Check |  | PASS | PASS | Pass | pass | Pass
Max OFS Offset |  | 7.9 V | 5.75 V | 8.1V | 8 | 8
95% OFS Offset |  | 7.505 V | 5.45 V | 7.7 | 7.6 | 7.6
Operating OFS Offset |  | 3.7525 V | 2.725 V | 3.85 | 3.8 | 3.8
Laser Output Power | 2.03 W | 1.94 W | 1.92 W | 1.93W | 1.94 W | 1.94W
After-Laser Rejected Power |  | 5.13 mW | 0.383 mW | 4.1mW | 4.11 mW | 3.96mW
AOM Input Power |  | 1.84 W | 1.86 W | 1.91W | 1.91 W | 1.88W
Max Diffracted Power |  | 1.55 W | 1.05 W | 1.60W | 1.60 W | 1.58 W
Un-Diffracted Power |  | 256 mW | 0.74 W | 160mW | 0.285 W | 155mW
AOM Diffraction Efficiency |  | 84.24% | 49% | 83.8% | 83.77% | 84.04%
After-AOM Rejected Power |  | 16.4 mW | 15.8 mW | 16.5mW | 15.5 mW | 13.9mW
TxPD Power |  | 6.39 mW | 4.05 mW | 14.2mW | 6.51mW | 13.7mW
OFSPD Power |  | 6.5 mW | 4.24mW | 6.68mw | 6.64 mW | 6.59mW
Outer Beam Power |  | 0.716 W | 0.25 W | 0.773W | 0.355 W | 0.747W
Inner Beam Power |  | 0.714 W | 0.248 W | 0.782W | 0.352 W | 0.744W
Output Beam Power Ratio |  | 0.9972 | 0.992 | 0.9885 | 0.99155 | 0.995
OFS Gain |  | 38.20 dB | 43 dB | 50.6 | 38.5 | 37.79
OFS Phase Margin |  | 55 degrees | 56 degrees | 61.7deg | 58 | 57.6

## LLO End X
Laser SN06 | Incoming laser power from T1300100 | X-END
Date | 2013-06-04
Laser Shutter Check |  | Alog: | 77828 | 59719 | 55574 | 49690 | 41046
Max OFS Offset |  |  | 29 July 2025 | 18 April 2022 | 20 April 2021 | 5 November 2019 | 8 October 2018
95% OFS Offset |  | Immediatly after the shutter | 1.89 W | 1.98 W | 1.80 W | 1.79 W | 1.80 W
Operating OFS Offset |  | Incident into AOM | 1.85 W | 1.87 W | 1.74 W | 1.74 W | 1.80 W
Laser Output Power | 2.13 W | Catching all beams out of AOM | 1.84 W | - | 1.54 W | - | -
Rejected Laser Power |  | AOM Diffraction Efficiency | - | 75% | 62% | 67% | 64%
AOM Input Power |  | Max OFS Offset | - | 8.2 V | 6.8 V | 7.14 V | -
Max Diffracted Power |  | OFS 95% | - | 7.79 V | 6.46 V | 6.78 V | -
Un-Diffracted Power |  | Operating OFS offset | - | 3.895 V | 3.23 V | 3.39 V | -
AOM Diffraction Efficiency |  | Max diffracted power | - | 1.40 W | 1.12 W | 1.17 W | 1.25 W
AOM Rejected Power |  | Undiffracted power | - | 0.60 W * | 0.43 W | 0.46 W | 0.45 W
TxPD Power
OFSPD Power
Outer Beam Power |  |  |  | Note: In this measurements there are two measurements recorded, one was done after tweaking current and temp. The one above is the AFTER the tweaking. |  |  | Note: This data was taken from the comment of the alog (41101).
Inner Beam Power
Output Beam Power Ratio
OFS Gain |  |  |  | * Before power tweaks.
OFS Phase Margin

## LLO End Y
Laser SN07 | Incoming laser power from T1300100 | Y-END
Date | 2013-06-04
Laser Shutter Check |  | Alog: | 79017 | 59884 | 55574 | 49505
Max OFS Offset |  |  | 14 October 2025 | 3 May 2022 | 20 April 2021 | 29 October 2019
95% OFS Offset |  | Immediatly after the shutter | 1.47 W | 1.63 W | 1.46 W | 1.42 W
Operating OFS Offset |  | Incident into AOM | 1.45 W | 1.58 W | 1.42 W | 1.39 W
Laser Output Power | 2.11 W | Catching all beams out of AOM |  | - | 0.95 W | 1.04 W
Rejected Laser Power |  | AOM Diffraction Efficiency | 60% | 65% | 56% | close to 80%
AOM Input Power |  | Max OFS Offset | 6.75 V | - | 6.0 V | 6.3 V
Max Diffracted Power |  | OFS 95% | 6.4 V | - | 5.7 V | 5.99 V
Un-Diffracted Power |  | Operating OFS offset |  | - | 2.85 V | 2.99 V
AOM Diffraction Efficiency |  | Max diffracted power | 0.871 W | 1.03 W | 0.80 W | 0.843 W
AOM Rejected Power |  | Undiffracted power |  | 0.15 W * | 0.14 W | 123 mW
TxPD Power
OFSPD Power |  |  |  | Note: In this measurements there are two measurements recorded, one was done after tweaking current and temp. The one above is the AFTER the tweaking.
Outer Beam Power
Inner Beam Power |  |  |  | * Before power tweaks.
Output Beam Power Ratio
OFS Gain
OFS Phase Margin

## LHO End X
Laser SN01 | Incoming laser power from T1300100
Date | 2013-06-04 | Aug. 3, 2018 | Mar. 19, 2019 | Jan-18-2024 | April-22-2025 | Jan-27-2026
Laser Shutter Check |  | PASS | PASS | Pass | Pass | Pass
Max OFS Offset |  | 9.9 V | 7.8 V | 7.7V | 7.7V | 7.75
95% OFS Offset |  | 9.405 V | 7.4 V | 7.3V | 7.315 | 7.36
Operating OFS Offset |  | 4.7025 V | 3.7 V | 3.85V | 3.6575 | 3.68
Laser Output Power | 2.05 W | 1.72 W | 1.75 W | 1.67W | 1.7 | 1.662
After-Laser Rejected Power |  | 56.5 mW | 54 mW | 36.6mW | 34.6mW | 35.7mW
AOM Input Power |  | 1.60 W | 1.69 W | 1.59W | 1.65 | 1.64
Max Diffracted Power |  | 0.967 W | 1.13 W | 1.1W | 1.14 | 1.1
Un-Diffracted Power |  | 0.359 W | 0.306 W | 266mW | 0.307W | 0.297
AOM Diffraction Efficiency |  | 60.44% | 67% | 69% | 0.69091 | 0.67073
After-AOM Rejected Power |  | 12.8 mW | 13.7 mW | 14.7mW | 15.2mW | 16.7mW
TxPD Power |  | 11.1 mW | 6.06 mW | 6.36mW | 11.9mW | 6.14mW
OFSPD Power |  | 5.7 mW | 6.56 mW | 6.36mW | 5.7mW | 6.58mW
Outer Beam Power |  | 228 mW | 0.257 W | 265mW | 0.536 | 512
Inner Beam Power |  | 228 mW | 0.255 W | 260mW | 0.526 | 501
Output Beam Power Ratio |  | 1 | 0.99222 | 0.9811 | 0.98134 | 0.97852
OFS Gain |  | 38.19 dB | 41 dB | 39.6db | 39.6 | 39.79
OFS Phase Margin |  | 60.9 degrees | 55 degrees | 43deg | 42.1 | 42.3
ALOG |  |  |  | 75484 |  | 88949

## LHO Pcal lab test setup
Laser SN05 | Incoming laser power from T1300100
Date | 2013-06-04 | Jan-17-2024 | June-26-2025
Laser Shutter Check |  | Pass | Fail
Max OFS Offset |  | 7.8V | 7.2V
95% OFS Offset |  | 7.4V | 6.84V
Operating OFS Offset |  | 3.7V | 3.42V
Laser Output Power |  | 1.98W | 1.97W
Rejected Laser Power |  | 15mW | 3.3mW
AOM Input Power |  | 1.92W | 1.92W
Max Diffracted Power |  | 1.32W | 1.23W
Un-Diffracted Power |  | 440mW (-1 order 90mW) | 0.500W
AOM Diffraction Efficiency |  | 69% | 64%
AOM Rejected Power |  | 21.3mW | 3.3mW
TxPD Power |  | 7.67mW | 1.4mW
OFSPD Power |  | 9.43mW | 1.73mW
Outer Beam Power |  | 318mw | 294mW
Inner Beam Power |  | 310mW | 292mW
Output Beam Power Ratio |  | 0.975 | 0.9932
OFS Gain |  | 7.39 | 7.50V
OFS Phase Margin |  | 43.6 | 43.7
ALOG |  | 75452